## Slide 1: ICU
Anesthesia Technologist BCS. Anesthesia. and IC diploma. Community health Karrar Nader AL-Taie
Anesthesia Technologist BCS. Anesthesia. and IC diploma. Community health Muneer Salman Hasan
L3-3

## Slide 2
Artificial Ventilation
Part 3

## Slide 3: Pressure mode
pressure is consistent with these modes, Volume is not
inspiration is terminated when a specific airway pressure has been reached.
inspiration stops when the preset inspiratory pressure is achieved
pressure mode delivers a selected gas pressure to the patient early in inspiration, and sustains the pressure throughout the inspiratory phase.

## Slide 4: Pressure modes
1-Pressure-controlled ventilation (PCV)
2- Pressure-support ventilation (PSV)
3- Invasive Continuous positive airway pressure (CPAP)
4- Noninvasive bi-level positive airway pressure ventilation (BiPAP)
5- High-frequency oscillatory ventilation (HFOV)
6-Airway pressure release ventilation (APRV)

## Slide 5: 1-Pressure-controlled ventilation (PCV)
the breathing gas flows under constant pressure into the lungs during the selected inspiratory time.
Normal peak inspiratory pressure (PIP) is 25-30 cm H2O.
(PIP) should be kept below 20 to 25 cm H2O whenever positive-pressure ventilation is required
The flow is highest at the beginning of inspiration (i.e. when the volume is lowest in the lungs).
Tidal volume varies with compliance and airway resistance and must be closely monitored.
PIP . I:E ratio/ Ti . RR must be selected.
Sedation and the use of neuromuscular blocking agents are frequently indicated, because any patient–ventilator asynchrony usually results in profound drops in the SaO2.

## Slide 8: Use PCV mode
1-ventilation of neonates
2- most complex ventilatory challenges, including severe acute respiratory distress syndrome (ARDS)
3- extremes of airflow obstruction (COPD)
4- difficult to wean patient.

## Slide 9
Advantages
1-Tidal volume variable with constant peak airway pressure
2-Full ventilatory support
3- Decreased mean airway pressure
4- Control frequency
Disadvantages
1- Requires sedation or paralysis
2- Ventilation does not change in response to clinical changing

## Slide 10: 2- Pressure-support ventilation (PSV)
the patient breathes spontaneously while the ventilator applies a predetermined amount of positive pressure to the airways upon inspiration.
Pressure support ventilation may be combined with SIMV or used alone for a spontaneously breathing patient.
In PSV mode, the inspired tidal volume and respiratory rate must be monitored closely to detect changes in lung compliance
It is a mode used primarily for weaning from mechanical ventilation.

## Slide 12
Advantages of PSV
1-Avoids patient-ventilator asynchrony
2-Patient comfortable-having full control over their ventilatory pattern and minute ventilation
3-Avoids breath stacking and auto PEEP
Disadvantages
1-Spontaneous mode -can't be used in heavily sedated, paralyzed or comatose patients
2- Respiratory muscle fatigue can develop if pressure support is set too low

## Slide 13: 3-Invasive CPAP
Constant positive airway pressure during spontaneous breathing
CPAP allows to observe the ability of the patient to breathe spontaneously while still on the ventilator.
CPAP can be used for intubated and nonintubated patients.
It may be used as a weaning mode

## Slide 14: 4- Noninvasive bi-level positive airway pressure ventilation (BiPAP)
BiPAP is a noninvasive form of mechanical ventilation provided by means of a nasal mask or nasal prongs, or a full-face mask.
The system allows the clinician to select two levels of positive-pressure support:
-An inspiratory pressure support level (referred to as IPAP)
-An expiratory pressure called EPAP (PEEP/CPAP level).

## Slide 15: Weaning:
the term weaning is used to describe the gradual process of decreasing ventilator support.
Methods of weaning:
1- T-piece trial,
2- Continuous Positive Airway Pressure (CPAP) weaning,
3- Synchronized Intermittent Mandatory Ventilation (SIMV) weaning,
4-Pressure Support Ventilation (PSV) weaning

## Slide 16: T-piece trial:
It consists of removing the patient from the ventilator and having him / her breathe spontaneously on a T-tube connected to oxygen source.
During T-piece weaning, periods of ventilator support are alternated with spontaneous breathing.
The goal is to progressively increase the time spent off the ventilator

## Slide 17: SIMV weaning
SIMV is the most common method of weaning.
It consists of gradually decreasing the number of breaths delivered by the ventilator to allow the patient to increase number of spontaneous breaths

## Slide 18: CPAP weaning:
When placed on CPAP, the patient does all the work of breathing without the aid of a back up rate or tidal volume.
No mandatory breaths are delivered in this mode . All ventilation is spontaneously initiated by the patient.
Weaning by gradual decrease in pressure value

## Slide 19: PSV weaning:
level of pressure support is gradually decreased based on the patient maintaining an adequate tidal volume (8 to 12 mL/kg) and a respiratory rate of less than 25 breaths/minute.
PSV weaning is indicated for: -
- Difficult to wean patients
- Small spontaneous tidal volume.

## Slide 20: Weaning readiness criteria:
1-Awake and alert o Hemodynamically stable,
2- Arterial blood gases (ABGs) normalized or at patient’s baseline
- PaCO2 acceptable
- PH of 7.35 – 7.45
- PaO2 > 60 mm Hg ,
- SaO2 >92%
- FIO2 ≤40% o
3- (PEEP) ≤ 5 cm H2O
4- F /RR< 25 / minute
5- Vt 5 ml / kg
6- VE 5- 10 L/m (f x Vt)
7- VC > 10- 15 ml / kg
8- Chest x-ray reviewed for correctable factors; treated as indicated,
9- Major electrolytes within normal range,
10- Hematocrit >25%,
11- Core temperature >36°C and

## Slide 21: Signs of weaning intolerance:
1-Diaphoresis
2-Dyspnea & Labored respiratory pattern
3- Increased anxiety, Restlessness, Decrease in level of consciousness
4- Dysrhythmia ,Increase or decrease in heart rate of > 20 beats /min.
5-Increase or decrease in blood pressure of > 20 mm Hg
6-RR greater than 35 breaths/minute
7-Vt ≤5 mL/kg, ME< 5 L/Min SPO2 <90%,
PaO2 < 60 mmHg, PH of < 7.35. Increase in PaCO2