_                                                                                                             Lecture 6 : Manual and automated  RED BLOOD CELL (RBC’s) COUNT

The RBC counting:  The count of red blood cells, is the number of red blood cells in mmᶟ of whole blood.

RBC count - Normal Values :

  - Men: 4.5-6.2 × 106/mm3

  - Women: 4.0-5.5 × 106/mm3

  - at birth: 4.0-6.0 × 106/mm3

  - 3 years – 10 years: 4.0-5.0 × 106/mm3

The  RBC counting  performed :

- Manually by  Heamocytometer  or

- The blood cells automatic  analyzer (Part of  CBC ;group of tests)

Importance  OF  RBCS  counting :

RBCs count  less or more than normal , It is a screen test anemia or   polycythemia

- Rbc count using  Haemocytometer :

Materials:

- Venous blood mixing with EDTA or capillary blood with heparin

- Improved  neubauer's  counting chamber with   Coverslips

- RBC  pipette and RBC  Diluting solution (Hayem's  Fluid ), the purpose of this fluid: its isotonic solution diluted blood, prevent lysis and blood sedimentation)

- Compound  Microscope

Principle: The blood is diluted 200 time with RBCs diluting fluid (1\200) in practice (20ul blood and 4 ml dilution).

Procedure:

- Draw blood by micropipette 20 ul.

- Mix the blood with diluting fluid 4ml.

- Mix the contents in glass tube for 2 min.

- Introducing the sample into the Neubauer chamber (10ul from mixture).

- To count the RBCs the microscope must be switched to 40X objective

- Count  RBCs from  R marked  5 squares

- Rbcs count in*10.000 cells of blood

Causes of decreased RBC count:

- Impaired red blood cell (RBC) production - Increased RBC destruction (hemolytic anemia's) - Blood loss - Fluid overload (hemodilution)

Causes of high red blood cell count:

- High altitudes , Hypoxia, Smoking, Dehydration

- Polycythemia vera (often a hereditary problem)

- Congestive heart disease, Certain lung diseases

- Erythropoietin doping by athletes to boost their performance

- Some types of anemia has high RBC count such as β-thalassemia

Physiological effect on Rbc count: -

1-age 2-sex 3-Activity 4-nutrition 5-pregnancy 6-brest feeding 7-psychological Emotions

2. Automated Red blood cell count:

Electronic counter is based on the principle of aperture impedance method.