## Slide 1
Dr. Ahmed Jamil Al- Taie
Laboratory: 3
First stage
Medical laboratories instruments

## Slide 2: White blood cells (Leukocytes)
-A type of blood cell that is made in the bone marrow
-WBCs (Leukocytes) are round shaped, nucleated, colorless cells. They are found throughout the body including the blood and lymphatic system.
-they are much fewer in number than red blood cells, accounting for about 1 percent of your blood.
-White blood cells are part of the body's immune system.
-They help the body fight infection and other diseases.

## Slide 3
-The White blood cells count is important in the diagnosis of disease
-Normal Value of Whit blood cells
Infants 15,000-20,000 cell/mm3
Children 4,500-13,500 cell/mm3
Adult 4,000-11,000 cell/mm3

## Slide 4: 1.Granulocytes A- basophil B- eosinophilic C- neutrophilic 2- A granulocytes A- monocytes B- lymphocytes

## Slide 5
1.Granulocytes

## Slide 6: 2. Agranulocytes

## Slide 7
. The Principle
-The blood is diluted 20 times with the WBCs diluting fluid .
-The leukocytes are then counted in 4 corners squares of the
improved Neubauer chamber.
Purpose
The WBCs count is to determine the number of circulating WBCs in the body

## Slide 8: Materials and instruments
Hemocytometer counting chamber (Neubauer chamber with coverslip).
WBCs counting pipette.
Rubber tube with a plastic mouth piece for drawing the fluid into a pipette.
Microscope
Diluting fluid.
Disposable blood lancet, sterile cotton, & 70% alcohol

## Slide 9: Diluting fluid (Turk’s solution )
This fluid is used for diluting the blood, it is composed of:
Glacial acetic acid (1.5 ml) (hemolyzes RBCs without affecting WBCs).
Gentian violet (1% solution) (1.5 ml) (stains the nuclei of WBC).
Distilled water to (97 ml).

## Slide 10: pipette
A--WBC pipette have white bead in the bulb and its mark (0.5, 1, 11)
B- While RBC pipette have red bead in the bulb and its mark (0.5, 1, 101)

## Slide 11: Procedure
1- Ensure that Haemocytometer chamber, coverslips and lenses of the microscope are clean, WBC pipette is dray.
WBC pipette
Haemocytometer
microscope

## Slide 12: 2. Sterilize the tip of the middle finger of hand
3. Suck the blood up to 0.5 marking of WBC pipette followed by the diluting fluid (Thomas solution) up to 11 marking without any air bubble.

## Slide 13
Mixing is done by rotating the pipette between the palms
After mixing the blood with diluting fluid, keep it aside for 3 minutes
Touch the tip of the pipette on the surface of the counting chamber in angle 45 degree.
Place the chamber on the stage of the microscope and allow 2 minutes for the cells to settle.
Scan the counting area with 10x objective lens then 40x objective

## Slide 14: Mixing of sample

## Slide 16: Calculation number of WBCs

## Slide 17: Calculation
The total number of WBCs/mm3 of
4 sequares = N × 50.
The total number of WBCs/mm3 of
2 sequares = N × 100.

## Slide 18: Increased WBCs In counts
Increased in numbers of WBC called (Leukocytosis)
it is due to: -
Bacterial infection, Inflammation, Parasites and Leukemia
Physiology leukocytosis as exercise
Pregnancy in the final month

## Slide 19: Decreased WBCs In counts
Decreased in numbers of WBC called (Leukopenia) it is due to:-
Viral infection (HIV, Hepatitis)
Nutritional deficiency such as (vitamin B12, folate deficiency)